## CSC_calc_0.6to0.8V
Paste data from a single Gamry Echem Analyst CV curve into cell A8. | [14 empty] | Calc. Diameter for Different GSA Values
This integration routine is intended for Iridium Oxide or other materials cycled between limits of -0.6 and 0.8 V with 10 mV step size. | [10 empty] | Sweep Rate | 50 | mV/s |  | Diameter (um) | Area (um^2) | Area (cm^2)
If you are using different parameters, you will have to adjust the length of the data and analysis columns in the spreadsheet. | [10 empty] | Resolution | 10 | mV |  | 11.2838 | 100 | 1e-06
Enter 1 or -1 for correct Gamry current sign convention: |  |  |  |  | 1 | 1 or -1 |  |  |  |  | Sampling | 0.2 | s |  | 25.2313 | 500 | 5e-06
If the Gamry convention is cathodic current = negative values, use " 1 " | [14 empty] | 35.6825 | 1000 | 1e-05
[11 empty] | Diam (mm) | 9.8693 | um |  | 50.4627 | 2000 | 2e-05
^^^^^^^^^^^^ | [10 empty] | Area | 76.4614 | um^2 |  | 61.8039 | 3000 | 3e-05
XY |  |  | Time | Pot | Charge | Running Total Charge | Running Total Charge Density | Charge Density | Cathodal Integral values | Anodal Integral values |  |  |  |  | 71.365 | 4000 | 4e-05
| CURVE3 (20250325_ZA_UM0004_A_E06_PBSair_CV50.DTA) |  | s | V | mC | mC | mC/cm2 | mC/cm2 | mC/cm2 | mC/cm2 | Area for Calc. | 7.6461e-07 | cm^2 |  | 79.7885 | 5000 | 5e-05
0.7901 | -1.057e-09 |  | 0 | 0.7901 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  |  | 112.8379 | 10000 | 0.0001
0.7801 | -1.25e-09 |  | 0.2 | 0.7801 | -2.307e-07 | -2.307e-07 | -0.30172 | 0.30172 | 0.30172 | 0 |  |  |  |  | Input Custom Area Value:
0.7701 | -1.288e-09 |  | 0.4 | 0.7701 | -2.538e-07 | -4.845e-07 | -0.63365 | 0.33193 | 0.33193 | 0 | EIC Method |  |  |  | 9.8693 | 76.5 | 7.65e-07
0.7601 | -1.286e-09 | Cathodic: | 0.6 | 0.7601 | -2.574e-07 | -7.419e-07 | -0.97029 | 0.33664 | 0.33664 | 0
0.7501 | -1.278e-09 | 54.9456 | 0.8 | 0.7501 | -2.564e-07 | -9.983e-07 | -1.3056 | 0.33533 | 0.33533 | 0 | CSCc | 54.9456 | mC/cm^2
0.7401 | -1.27e-09 |  | 1 | 0.7401 | -2.548e-07 | -1.2531e-06 | -1.6389 | 0.33324 | 0.33324 | 0 | Charge | 42.0122 | nC
0.7301 | -1.268e-09 | HMRI: | 1.2 | 0.7301 | -2.538e-07 | -1.5069e-06 | -1.9708 | 0.33193 | 0.33193 | 0
0.7201 | -1.265e-09 | 46.4309 | 1.4 | 0.7201 | -2.533e-07 | -1.7602e-06 | -2.3021 | 0.33128 | 0.33128 | 0 | CSCa | 37.9161 | mC/cm^2
0.7102 | -1.258e-09 |  | 1.6 | 0.7102 | -2.523e-07 | -2.0125e-06 | -2.632 | 0.32997 | 0.32997 | 0 | Charge | 28.9912 | nC
0.7001 | -1.26e-09 |  | 1.8 | 0.7001 | -2.518e-07 | -2.2643e-06 | -2.9614 | 0.32932 | 0.32932 | 0
0.6901 | -1.274e-09 |  | 2 | 0.6901 | -2.534e-07 | -2.5177e-06 | -3.2928 | 0.33141 | 0.33141 | 0 | HMRI Method
0.6802 | -1.291e-09 |  | 2.2 | 0.6802 | -2.565e-07 | -2.7742e-06 | -3.6282 | 0.33546 | 0.33546 | 0 | Total Absolute Charge | 92.8617 | mC/cm^2
0.6701 | -1.309e-09 |  | 2.4 | 0.6701 | -2.6e-07 | -3.0342e-06 | -3.9683 | 0.34004 | 0.34004 | 0 | Divided by 2…
0.6601 | -1.324e-09 |  | 2.6 | 0.6601 | -2.633e-07 | -3.2975e-06 | -4.3126 | 0.34436 | 0.34436 | 0 | CSCc | 46.4309 | mC/cm^2
0.6501 | -1.345e-09 |  | 2.8 | 0.6501 | -2.669e-07 | -3.5644e-06 | -4.6617 | 0.34906 | 0.34906 | 0 | Charge | 35.5017 | nC
0.6401 | -1.365e-09 |  | 3 | 0.6401 | -2.71e-07 | -3.8354e-06 | -5.0161 | 0.35443 | 0.35443 | 0
0.6301 | -1.392e-09 |  | 3.2 | 0.6301 | -2.757e-07 | -4.1111e-06 | -5.3767 | 0.36057 | 0.36057 | 0
0.6201 | -1.417e-09 |  | 3.4 | 0.6201 | -2.809e-07 | -4.392e-06 | -5.7441 | 0.36737 | 0.36737 | 0 |  | CSCc | CSCa | CSCh
0.6101 | -1.445e-09 |  | 3.6 | 0.6101 | -2.862e-07 | -4.6782e-06 | -6.1184 | 0.37431 | 0.37431 | 0 |  | 54.9456 | 37.9161 | 46.4309 | mC/cm^2
0.6001 | -1.473e-09 |  | 3.8 | 0.6001 | -2.918e-07 | -4.97e-06 | -6.5 | 0.38163 | 0.38163 | 0
0.5901 | -1.502e-09 |  | 4 | 0.5901 | -2.975e-07 | -5.2675e-06 | -6.8891 | 0.38909 | 0.38909 | 0 |  | Qc | Qa | Qh
0.5801 | -1.53e-09 |  | 4.2 | 0.5801 | -3.032e-07 | -5.5707e-06 | -7.2856 | 0.39654 | 0.39654 | 0 |  | 42.0122 | 28.9912 | 35.5017 | nC
0.5701 | -1.551e-09 |  | 4.4 | 0.5701 | -3.081e-07 | -5.8788e-06 | -7.6886 | 0.40295 | 0.40295 | 0
0.5601 | -1.562e-09 |  | 4.6 | 0.5601 | -3.113e-07 | -6.1901e-06 | -8.0957 | 0.40713 | 0.40713 | 0
0.5501 | -1.575e-09 |  | 4.8 | 0.5501 | -3.137e-07 | -6.5038e-06 | -8.506 | 0.41027 | 0.41027 | 0
0.5401 | -1.584e-09 |  | 5 | 0.5401 | -3.159e-07 | -6.8197e-06 | -8.9191 | 0.41315 | 0.41315 | 0
0.5301 | -1.579e-09 |  | 5.2 | 0.5301 | -3.163e-07 | -7.136e-06 | -9.3328 | 0.41367 | 0.41367 | 0
0.5201 | -1.571e-09 |  | 5.4 | 0.5201 | -3.15e-07 | -7.451e-06 | -9.7448 | 0.41197 | 0.41197 | 0
0.5101 | -1.552e-09 |  | 5.6 | 0.5101 | -3.123e-07 | -7.7633e-06 | -10.1532 | 0.40844 | 0.40844 | 0
0.5001 | -1.535e-09 |  | 5.8 | 0.5001 | -3.087e-07 | -8.072e-06 | -10.557 | 0.40373 | 0.40373 | 0
0.4901 | -1.514e-09 |  | 6 | 0.4901 | -3.049e-07 | -8.3769e-06 | -10.9557 | 0.39876 | 0.39876 | 0
0.4801 | -1.488e-09 |  | 6.2 | 0.4801 | -3.002e-07 | -8.6771e-06 | -11.3483 | 0.39262 | 0.39262 | 0
0.4701 | -1.445e-09 |  | 6.4 | 0.4701 | -2.933e-07 | -8.9704e-06 | -11.7319 | 0.38359 | 0.38359 | 0
0.4601 | -1.409e-09 |  | 6.6 | 0.4601 | -2.854e-07 | -9.2558e-06 | -12.1052 | 0.37326 | 0.37326 | 0
0.4501 | -1.374e-09 |  | 6.8 | 0.4501 | -2.783e-07 | -9.5341e-06 | -12.4692 | 0.36397 | 0.36397 | 0
0.4401 | -1.337e-09 |  | 7 | 0.4401 | -2.711e-07 | -9.8052e-06 | -12.8237 | 0.35456 | 0.35456 | 0
0.4301 | -1.295e-09 |  | 7.2 | 0.4301 | -2.632e-07 | -1.0068e-05 | -13.1679 | 0.34423 | 0.34423 | 0
0.4201 | -1.26e-09 |  | 7.4 | 0.4201 | -2.555e-07 | -1.0324e-05 | -13.5021 | 0.33416 | 0.33416 | 0
0.4101 | -1.235e-09 |  | 7.6 | 0.4101 | -2.495e-07 | -1.0573e-05 | -13.8284 | 0.32631 | 0.32631 | 0
0.4001 | -1.204e-09 |  | 7.8 | 0.4001 | -2.439e-07 | -1.0817e-05 | -14.1474 | 0.31898 | 0.31898 | 0
0.3901 | -1.189e-09 |  | 8 | 0.3901 | -2.393e-07 | -1.1057e-05 | -14.4604 | 0.31297 | 0.31297 | 0
0.3801 | -1.175e-09 |  | 8.2 | 0.3801 | -2.364e-07 | -1.1293e-05 | -14.7695 | 0.30918 | 0.30918 | 0
0.3701 | -1.171e-09 |  | 8.4 | 0.3701 | -2.346e-07 | -1.1528e-05 | -15.0764 | 0.30682 | 0.30682 | 0
0.3601 | -1.178e-09 |  | 8.6 | 0.3601 | -2.349e-07 | -1.1762e-05 | -15.3836 | 0.30721 | 0.30721 | 0
0.3501 | -1.198e-09 |  | 8.8 | 0.3501 | -2.376e-07 | -1.2e-05 | -15.6943 | 0.31074 | 0.31074 | 0
0.3401 | -1.234e-09 |  | 9 | 0.3401 | -2.432e-07 | -1.2243e-05 | -16.0124 | 0.31807 | 0.31807 | 0
0.3301 | -1.28e-09 |  | 9.2 | 0.3301 | -2.514e-07 | -1.2495e-05 | -16.3412 | 0.32879 | 0.32879 | 0
0.3201 | -1.341e-09 |  | 9.4 | 0.3201 | -2.621e-07 | -1.2757e-05 | -16.684 | 0.34279 | 0.34279 | 0
0.3101 | -1.406e-09 |  | 9.6 | 0.3101 | -2.747e-07 | -1.3031e-05 | -17.0432 | 0.35927 | 0.35927 | 0
0.3001 | -1.484e-09 |  | 9.8 | 0.3001 | -2.89e-07 | -1.332e-05 | -17.4212 | 0.37797 | 0.37797 | 0
0.2901 | -1.568e-09 |  | 10 | 0.2901 | -3.052e-07 | -1.3626e-05 | -17.8204 | 0.39916 | 0.39916 | 0
0.28 | -1.673e-09 |  | 10.2 | 0.28 | -3.241e-07 | -1.395e-05 | -18.2442 | 0.42387 | 0.42387 | 0
0.2701 | -1.762e-09 |  | 10.4 | 0.2701 | -3.435e-07 | -1.4293e-05 | -18.6935 | 0.44925 | 0.44925 | 0
0.2601 | -1.863e-09 |  | 10.6 | 0.2601 | -3.625e-07 | -1.4656e-05 | -19.1676 | 0.4741 | 0.4741 | 0
0.2501 | -1.961e-09 |  | 10.8 | 0.2501 | -3.824e-07 | -1.5038e-05 | -19.6677 | 0.50012 | 0.50012 | 0
0.2401 | -2.071e-09 |  | 11 | 0.2401 | -4.032e-07 | -1.5441e-05 | -20.195 | 0.52732 | 0.52732 | 0
0.2301 | -2.162e-09 |  | 11.2 | 0.2301 | -4.233e-07 | -1.5865e-05 | -20.7486 | 0.55361 | 0.55361 | 0
0.2201 | -2.242e-09 |  | 11.4 | 0.2201 | -4.404e-07 | -1.6305e-05 | -21.3246 | 0.57598 | 0.57598 | 0
0.2101 | -2.325e-09 |  | 11.6 | 0.2101 | -4.567e-07 | -1.6762e-05 | -21.9219 | 0.59729 | 0.59729 | 0
0.2001 | -2.391e-09 |  | 11.8 | 0.2001 | -4.716e-07 | -1.7233e-05 | -22.5387 | 0.61678 | 0.61678 | 0
0.1901 | -2.451e-09 |  | 12 | 0.1901 | -4.842e-07 | -1.7718e-05 | -23.1719 | 0.63326 | 0.63326 | 0
0.1801 | -2.5e-09 |  | 12.2 | 0.1801 | -4.951e-07 | -1.8213e-05 | -23.8195 | 0.64752 | 0.64752 | 0
0.1701 | -2.525e-09 |  | 12.4 | 0.1701 | -5.025e-07 | -1.8715e-05 | -24.4767 | 0.65719 | 0.65719 | 0
0.1601 | -2.543e-09 |  | 12.6 | 0.1601 | -5.068e-07 | -1.9222e-05 | -25.1395 | 0.66282 | 0.66282 | 0
0.1501 | -2.54e-09 |  | 12.8 | 0.1501 | -5.083e-07 | -1.973e-05 | -25.8043 | 0.66478 | 0.66478 | 0
0.1401 | -2.529e-09 |  | 13 | 0.1401 | -5.069e-07 | -2.0237e-05 | -26.4672 | 0.66295 | 0.66295 | 0
0.1301 | -2.491e-09 |  | 13.2 | 0.1301 | -5.02e-07 | -2.0739e-05 | -27.1237 | 0.65654 | 0.65654 | 0
0.1201 | -2.433e-09 |  | 13.4 | 0.1201 | -4.924e-07 | -2.1232e-05 | -27.7677 | 0.64398 | 0.64398 | 0
0.1101 | -2.361e-09 |  | 13.6 | 0.1101 | -4.794e-07 | -2.1711e-05 | -28.3947 | 0.62698 | 0.62698 | 0
0.1001 | -2.279e-09 |  | 13.8 | 0.1001 | -4.64e-07 | -2.2175e-05 | -29.0016 | 0.60684 | 0.60684 | 0
0.09004 | -2.18e-09 |  | 14 | 0.09004 | -4.459e-07 | -2.2621e-05 | -29.5847 | 0.58317 | 0.58317 | 0
0.08005 | -2.064e-09 |  | 14.2 | 0.08005 | -4.244e-07 | -2.3045e-05 | -30.1398 | 0.55505 | 0.55505 | 0
0.07007 | -1.944e-09 |  | 14.4 | 0.07007 | -4.008e-07 | -2.3446e-05 | -30.664 | 0.52419 | 0.52419 | 0
0.06003 | -1.82e-09 |  | 14.6 | 0.06003 | -3.764e-07 | -2.3822e-05 | -31.1562 | 0.49227 | 0.49227 | 0
0.05005 | -1.696e-09 |  | 14.8 | 0.05005 | -3.516e-07 | -2.4174e-05 | -31.6161 | 0.45984 | 0.45984 | 0
0.04004 | -1.582e-09 |  | 15 | 0.04004 | -3.278e-07 | -2.4502e-05 | -32.0448 | 0.42871 | 0.42871 | 0
0.03005 | -1.472e-09 |  | 15.2 | 0.03005 | -3.054e-07 | -2.4807e-05 | -32.4442 | 0.39942 | 0.39942 | 0
0.02005 | -1.377e-09 |  | 15.4 | 0.02005 | -2.849e-07 | -2.5092e-05 | -32.8168 | 0.37261 | 0.37261 | 0
0.01004 | -1.288e-09 |  | 15.6 | 0.01004 | -2.665e-07 | -2.5359e-05 | -33.1654 | 0.34854 | 0.34854 | 0
6.017e-05 | -1.199e-09 |  | 15.8 | 6.017e-05 | -2.487e-07 | -2.5607e-05 | -33.4906 | 0.32526 | 0.32526 | 0
-0.009954 | -1.128e-09 |  | 16 | -0.009954 | -2.327e-07 | -2.584e-05 | -33.795 | 0.30434 | 0.30434 | 0
-0.01996 | -1.071e-09 |  | 16.2 | -0.01996 | -2.199e-07 | -2.606e-05 | -34.0825 | 0.2876 | 0.2876 | 0
-0.02997 | -1.015e-09 |  | 16.4 | -0.02997 | -2.086e-07 | -2.6269e-05 | -34.3554 | 0.27282 | 0.27282 | 0
-0.03997 | -9.698e-10 |  | 16.6 | -0.03997 | -1.9848e-07 | -2.6467e-05 | -34.6149 | 0.25958 | 0.25958 | 0
-0.04994 | -9.234e-10 |  | 16.8 | -0.04994 | -1.8932e-07 | -2.6656e-05 | -34.8625 | 0.2476 | 0.2476 | 0
-0.05996 | -8.907e-10 |  | 17 | -0.05996 | -1.8141e-07 | -2.6838e-05 | -35.0998 | 0.23726 | 0.23726 | 0
-0.06994 | -8.563e-10 |  | 17.2 | -0.06994 | -1.747e-07 | -2.7013e-05 | -35.3283 | 0.22848 | 0.22848 | 0
-0.07996 | -8.327e-10 |  | 17.4 | -0.07996 | -1.689e-07 | -2.7181e-05 | -35.5492 | 0.2209 | 0.2209 | 0
-0.08994 | -8.156e-10 |  | 17.6 | -0.08994 | -1.6483e-07 | -2.7346e-05 | -35.7648 | 0.21557 | 0.21557 | 0
-0.09994 | -7.984e-10 |  | 17.8 | -0.09994 | -1.614e-07 | -2.7508e-05 | -35.9758 | 0.21109 | 0.21109 | 0
-0.11 | -7.808e-10 |  | 18 | -0.11 | -1.5792e-07 | -2.7666e-05 | -36.1824 | 0.20654 | 0.20654 | 0
-0.12 | -7.72e-10 |  | 18.2 | -0.12 | -1.5528e-07 | -2.7821e-05 | -36.3855 | 0.20308 | 0.20308 | 0
-0.13 | -7.631e-10 |  | 18.4 | -0.13 | -1.5351e-07 | -2.7974e-05 | -36.5862 | 0.20077 | 0.20077 | 0
-0.14 | -7.583e-10 |  | 18.6 | -0.14 | -1.5214e-07 | -2.8126e-05 | -36.7852 | 0.19898 | 0.19898 | 0
-0.15 | -7.512e-10 |  | 18.8 | -0.15 | -1.5095e-07 | -2.8277e-05 | -36.9826 | 0.19742 | 0.19742 | 0
-0.16 | -7.489e-10 |  | 19 | -0.16 | -1.5001e-07 | -2.8427e-05 | -37.1788 | 0.19619 | 0.19619 | 0
-0.17 | -7.458e-10 |  | 19.2 | -0.17 | -1.4947e-07 | -2.8577e-05 | -37.3743 | 0.19548 | 0.19548 | 0
-0.18 | -7.419e-10 |  | 19.4 | -0.18 | -1.4877e-07 | -2.8726e-05 | -37.5689 | 0.19457 | 0.19457 | 0
-0.19 | -7.45e-10 |  | 19.6 | -0.19 | -1.4869e-07 | -2.8874e-05 | -37.7633 | 0.19446 | 0.19446 | 0
-0.2 | -7.481e-10 |  | 19.8 | -0.2 | -1.4931e-07 | -2.9024e-05 | -37.9586 | 0.19527 | 0.19527 | 0
-0.21 | -7.527e-10 |  | 20 | -0.21 | -1.5008e-07 | -2.9174e-05 | -38.1549 | 0.19628 | 0.19628 | 0
-0.22 | -7.548e-10 |  | 20.2 | -0.22 | -1.5075e-07 | -2.9325e-05 | -38.352 | 0.19716 | 0.19716 | 0
-0.2299 | -7.608e-10 |  | 20.4 | -0.2299 | -1.5156e-07 | -2.9476e-05 | -38.5503 | 0.19822 | 0.19822 | 0
-0.24 | -7.737e-10 |  | 20.6 | -0.24 | -1.5345e-07 | -2.963e-05 | -38.751 | 0.20069 | 0.20069 | 0
-0.25 | -7.8e-10 |  | 20.8 | -0.25 | -1.5537e-07 | -2.9785e-05 | -38.9542 | 0.2032 | 0.2032 | 0
-0.26 | -7.864e-10 |  | 21 | -0.26 | -1.5664e-07 | -2.9942e-05 | -39.159 | 0.20486 | 0.20486 | 0
-0.27 | -7.972e-10 |  | 21.2 | -0.27 | -1.5836e-07 | -3.01e-05 | -39.3661 | 0.20711 | 0.20711 | 0
-0.2799 | -8.148e-10 |  | 21.4 | -0.2799 | -1.612e-07 | -3.0261e-05 | -39.577 | 0.21083 | 0.21083 | 0
-0.2899 | -8.343e-10 |  | 21.6 | -0.2899 | -1.6491e-07 | -3.0426e-05 | -39.7926 | 0.21568 | 0.21568 | 0
-0.3 | -8.553e-10 |  | 21.8 | -0.3 | -1.6896e-07 | -3.0595e-05 | -40.0136 | 0.22097 | 0.22097 | 0
-0.31 | -8.707e-10 |  | 22 | -0.31 | -1.726e-07 | -3.0768e-05 | -40.2393 | 0.22573 | 0.22573 | 0
-0.32 | -8.91e-10 |  | 22.2 | -0.32 | -1.7617e-07 | -3.0944e-05 | -40.4697 | 0.2304 | 0.2304 | 0
-0.33 | -9.114e-10 |  | 22.4 | -0.33 | -1.8024e-07 | -3.1124e-05 | -40.7055 | 0.23573 | 0.23573 | 0
-0.34 | -9.355e-10 |  | 22.6 | -0.34 | -1.8469e-07 | -3.1309e-05 | -40.947 | 0.24155 | 0.24155 | 0
-0.35 | -9.553e-10 |  | 22.8 | -0.35 | -1.8908e-07 | -3.1498e-05 | -41.1943 | 0.24729 | 0.24729 | 0
-0.36 | -9.794e-10 |  | 23 | -0.36 | -1.9347e-07 | -3.1691e-05 | -41.4473 | 0.25303 | 0.25303 | 0
-0.37 | -1.004e-09 |  | 23.2 | -0.37 | -1.9834e-07 | -3.189e-05 | -41.7067 | 0.2594 | 0.2594 | 0
-0.38 | -1.029e-09 |  | 23.4 | -0.38 | -2.033e-07 | -3.2093e-05 | -41.9726 | 0.26589 | 0.26589 | 0
-0.39 | -1.047e-09 |  | 23.6 | -0.39 | -2.076e-07 | -3.23e-05 | -42.2441 | 0.27151 | 0.27151 | 0
-0.4 | -1.06e-09 |  | 23.8 | -0.4 | -2.107e-07 | -3.2511e-05 | -42.5197 | 0.27556 | 0.27556 | 0
-0.41 | -1.077e-09 |  | 24 | -0.41 | -2.137e-07 | -3.2725e-05 | -42.7992 | 0.27949 | 0.27949 | 0
-0.42 | -1.092e-09 |  | 24.2 | -0.42 | -2.169e-07 | -3.2942e-05 | -43.0829 | 0.28367 | 0.28367 | 0
-0.43 | -1.102e-09 |  | 24.4 | -0.43 | -2.194e-07 | -3.3161e-05 | -43.3698 | 0.28694 | 0.28694 | 0
-0.44 | -1.107e-09 |  | 24.6 | -0.44 | -2.209e-07 | -3.3382e-05 | -43.6587 | 0.2889 | 0.2889 | 0
-0.45 | -1.105e-09 |  | 24.8 | -0.45 | -2.212e-07 | -3.3603e-05 | -43.948 | 0.2893 | 0.2893 | 0
-0.46 | -1.105e-09 |  | 25 | -0.46 | -2.21e-07 | -3.3824e-05 | -44.237 | 0.28903 | 0.28903 | 0
-0.47 | -1.104e-09 |  | 25.2 | -0.47 | -2.209e-07 | -3.4045e-05 | -44.5259 | 0.2889 | 0.2889 | 0
-0.48 | -1.104e-09 |  | 25.4 | -0.48 | -2.208e-07 | -3.4266e-05 | -44.8147 | 0.28877 | 0.28877 | 0
-0.49 | -1.105e-09 |  | 25.6 | -0.49 | -2.209e-07 | -3.4487e-05 | -45.1036 | 0.2889 | 0.2889 | 0
-0.5 | -1.106e-09 |  | 25.8 | -0.5 | -2.211e-07 | -3.4708e-05 | -45.3928 | 0.28917 | 0.28917 | 0
-0.51 | -1.107e-09 |  | 26 | -0.51 | -2.213e-07 | -3.4929e-05 | -45.6822 | 0.28943 | 0.28943 | 0
-0.52 | -1.114e-09 |  | 26.2 | -0.52 | -2.221e-07 | -3.5151e-05 | -45.9727 | 0.29047 | 0.29047 | 0
-0.53 | -1.127e-09 |  | 26.4 | -0.53 | -2.241e-07 | -3.5375e-05 | -46.2658 | 0.29309 | 0.29309 | 0
-0.54 | -1.149e-09 |  | 26.6 | -0.54 | -2.276e-07 | -3.5603e-05 | -46.5634 | 0.29767 | 0.29767 | 0
-0.55 | -1.201e-09 |  | 26.8 | -0.55 | -2.35e-07 | -3.5838e-05 | -46.8708 | 0.30734 | 0.30734 | 0
-0.56 | -1.282e-09 |  | 27 | -0.56 | -2.483e-07 | -3.6086e-05 | -47.1955 | 0.32474 | 0.32474 | 0
-0.57 | -1.417e-09 |  | 27.2 | -0.57 | -2.699e-07 | -3.6356e-05 | -47.5485 | 0.35299 | 0.35299 | 0
-0.58 | -1.603e-09 |  | 27.4 | -0.58 | -3.02e-07 | -3.6658e-05 | -47.9435 | 0.39497 | 0.39497 | 0
-0.59 | -1.837e-09 |  | 27.6 | -0.59 | -3.44e-07 | -3.7002e-05 | -48.3934 | 0.4499 | 0.4499 | 0
-0.6 | -2.12e-09 |  | 27.8 | -0.6 | -3.957e-07 | -3.7398e-05 | -48.9109 | 0.51752 | 0.51752 | 0
-0.59 | -1.585e-09 |  | 28 | -0.59 | -3.705e-07 | -3.7768e-05 | -49.3954 | 0.48456 | 0.48456 | 0
-0.58 | -1.312e-09 |  | 28.2 | -0.58 | -2.897e-07 | -3.8058e-05 | -49.7743 | 0.37888 | 0.37888 | 0
-0.57 | -1.103e-09 |  | 28.4 | -0.57 | -2.415e-07 | -3.83e-05 | -50.0902 | 0.31585 | 0.31585 | 0
-0.56 | -9.533e-10 |  | 28.6 | -0.56 | -2.0563e-07 | -3.8505e-05 | -50.3591 | 0.26893 | 0.26893 | 0
-0.55 | -8.509e-10 |  | 28.8 | -0.55 | -1.8042e-07 | -3.8686e-05 | -50.5951 | 0.23596 | 0.23596 | 0
-0.54 | -7.821e-10 |  | 29 | -0.54 | -1.633e-07 | -3.8849e-05 | -50.8086 | 0.21357 | 0.21357 | 0
-0.53 | -7.346e-10 |  | 29.2 | -0.53 | -1.5167e-07 | -3.9001e-05 | -51.007 | 0.19836 | 0.19836 | 0
-0.52 | -6.982e-10 |  | 29.4 | -0.52 | -1.4328e-07 | -3.9144e-05 | -51.1944 | 0.18739 | 0.18739 | 0
-0.51 | -6.732e-10 |  | 29.6 | -0.51 | -1.3714e-07 | -3.9281e-05 | -51.3738 | 0.17936 | 0.17936 | 0
-0.5001 | -6.464e-10 |  | 29.8 | -0.5001 | -1.3196e-07 | -3.9413e-05 | -51.5463 | 0.17258 | 0.17258 | 0
-0.49 | -6.239e-10 |  | 30 | -0.49 | -1.2703e-07 | -3.954e-05 | -51.7125 | 0.16614 | 0.16614 | 0
-0.48 | -6.051e-10 |  | 30.2 | -0.48 | -1.229e-07 | -3.9663e-05 | -51.8732 | 0.16073 | 0.16073 | 0
-0.47 | -5.917e-10 |  | 30.4 | -0.47 | -1.1968e-07 | -3.9783e-05 | -52.0297 | 0.15652 | 0.15652 | 0
-0.46 | -5.732e-10 |  | 30.6 | -0.46 | -1.1649e-07 | -3.9899e-05 | -52.1821 | 0.15235 | 0.15235 | 0
-0.45 | -5.6e-10 |  | 30.8 | -0.45 | -1.1332e-07 | -4.0012e-05 | -52.3303 | 0.14821 | 0.14821 | 0
-0.44 | -5.435e-10 |  | 31 | -0.44 | -1.1035e-07 | -4.0123e-05 | -52.4746 | 0.14432 | 0.14432 | 0
-0.43 | -5.292e-10 |  | 31.2 | -0.43 | -1.0727e-07 | -4.023e-05 | -52.6149 | 0.14029 | 0.14029 | 0
-0.42 | -5.146e-10 |  | 31.4 | -0.42 | -1.0438e-07 | -4.0334e-05 | -52.7514 | 0.13651 | 0.13651 | 0
-0.41 | -4.913e-10 |  | 31.6 | -0.41 | -1.0059e-07 | -4.0435e-05 | -52.883 | 0.13156 | 0.13156 | 0
-0.4 | -4.67e-10 |  | 31.8 | -0.4 | -9.583e-08 | -4.0531e-05 | -53.0083 | 0.12533 | 0.12533 | 0
-0.39 | -4.44e-10 |  | 32 | -0.39 | -9.11e-08 | -4.0622e-05 | -53.1274 | 0.11915 | 0.11915 | 0
-0.38 | -4.336e-10 |  | 32.2 | -0.38 | -8.776e-08 | -4.071e-05 | -53.2422 | 0.11478 | 0.11478 | 0
-0.37 | -4.185e-10 |  | 32.4 | -0.37 | -8.521e-08 | -4.0795e-05 | -53.3537 | 0.11144 | 0.11144 | 0
-0.36 | -4.071e-10 |  | 32.6 | -0.36 | -8.256e-08 | -4.0878e-05 | -53.4616 | 0.10798 | 0.10798 | 0
-0.35 | -3.94e-10 |  | 32.8 | -0.35 | -8.011e-08 | -4.0958e-05 | -53.5664 | 0.10477 | 0.10477 | 0
-0.34 | -3.854e-10 |  | 33 | -0.34 | -7.794e-08 | -4.1036e-05 | -53.6683 | 0.10193 | 0.10193 | 0
-0.33 | -3.768e-10 |  | 33.2 | -0.33 | -7.622e-08 | -4.1112e-05 | -53.768 | 0.099684 | 0.099684 | 0
-0.32 | -3.615e-10 |  | 33.4 | -0.32 | -7.383e-08 | -4.1186e-05 | -53.8646 | 0.096559 | 0.096559 | 0
-0.31 | -3.526e-10 |  | 33.6 | -0.31 | -7.141e-08 | -4.1257e-05 | -53.958 | 0.093394 | 0.093394 | 0
-0.3 | -3.446e-10 |  | 33.8 | -0.3 | -6.972e-08 | -4.1327e-05 | -54.0492 | 0.091183 | 0.091183 | 0
-0.29 | -3.315e-10 |  | 34 | -0.29 | -6.761e-08 | -4.1394e-05 | -54.1376 | 0.088424 | 0.088424 | 0
-0.28 | -3.2e-10 |  | 34.2 | -0.28 | -6.515e-08 | -4.146e-05 | -54.2228 | 0.085206 | 0.085206 | 0
-0.27 | -3.078e-10 |  | 34.4 | -0.27 | -6.278e-08 | -4.1522e-05 | -54.3049 | 0.082107 | 0.082107 | 0
-0.26 | -2.96e-10 |  | 34.6 | -0.26 | -6.038e-08 | -4.1583e-05 | -54.3839 | 0.078968 | 0.078968 | 0
-0.25 | -2.812e-10 |  | 34.8 | -0.25 | -5.772e-08 | -4.164e-05 | -54.4594 | 0.075489 | 0.075489 | 0
-0.24 | -2.67e-10 |  | 35 | -0.24 | -5.482e-08 | -4.1695e-05 | -54.5311 | 0.071696 | 0.071696 | 0
-0.23 | -2.482e-10 |  | 35.2 | -0.23 | -5.152e-08 | -4.1747e-05 | -54.5984 | 0.06738 | 0.06738 | 0
-0.22 | -2.277e-10 |  | 35.4 | -0.22 | -4.759e-08 | -4.1794e-05 | -54.6607 | 0.062241 | 0.062241 | 0
-0.21 | -2.044e-10 |  | 35.6 | -0.21 | -4.321e-08 | -4.1838e-05 | -54.7172 | 0.056512 | 0.056512 | 0
-0.2 | -1.771e-10 |  | 35.8 | -0.2 | -3.815e-08 | -4.1876e-05 | -54.7671 | 0.049894 | 0.049894 | 0
-0.19 | -1.577e-10 |  | 36 | -0.19 | -3.348e-08 | -4.1909e-05 | -54.8109 | 0.043787 | 0.043787 | 0
-0.18 | -1.343e-10 |  | 36.2 | -0.18 | -2.92e-08 | -4.1938e-05 | -54.8491 | 0.038189 | 0.038189 | 0
-0.17 | -1.145e-10 |  | 36.4 | -0.17 | -2.488e-08 | -4.1963e-05 | -54.8816 | 0.032539 | 0.032539 | 0
-0.16 | -8.929e-11 |  | 36.6 | -0.16 | -2.0379e-08 | -4.1984e-05 | -54.9083 | 0.026653 | 0.026653 | 0
-0.15 | -6.454e-11 |  | 36.8 | -0.15 | -1.5383e-08 | -4.1999e-05 | -54.9284 | 0.020119 | 0.020119 | 0
-0.14 | -3.168e-11 |  | 37 | -0.14 | -9.622e-09 | -4.2009e-05 | -54.941 | 0.012584 | 0.012584 | 0
-0.13 | -4.048e-12 |  | 37.2 | -0.13 | -3.5728e-09 | -4.2012e-05 | -54.9456 | 0.0046727 | 0.0046727 | 0
-0.12 | 2.939e-11 |  | 37.4 | -0.12 | 2.5342e-09 | -4.201e-05 | -54.9423 | 0.0033144 | 0 | 0.0033144
-0.11 | 6.15e-11 |  | 37.6 | -0.11 | 9.089e-09 | -4.2001e-05 | -54.9304 | 0.011887 | 0 | 0.011887
-0.09997 | 1.078e-10 |  | 37.8 | -0.09997 | 1.693e-08 | -4.1984e-05 | -54.9083 | 0.022142 | 0 | 0.022142
-0.08996 | 1.528e-10 |  | 38 | -0.08996 | 2.606e-08 | -4.1958e-05 | -54.8742 | 0.034083 | 0 | 0.034083
-0.07997 | 2.002e-10 |  | 38.2 | -0.07997 | 3.53e-08 | -4.1922e-05 | -54.828 | 0.046167 | 0 | 0.046167
-0.06997 | 2.611e-10 |  | 38.4 | -0.06997 | 4.613e-08 | -4.1876e-05 | -54.7677 | 0.060331 | 0 | 0.060331
-0.05997 | 3.171e-10 |  | 38.6 | -0.05997 | 5.782e-08 | -4.1818e-05 | -54.6921 | 0.07562 | 0 | 0.07562
-0.04996 | 3.845e-10 |  | 38.8 | -0.04996 | 7.016e-08 | -4.1748e-05 | -54.6003 | 0.091759 | 0 | 0.091759
-0.04 | 4.553e-10 |  | 39 | -0.04 | 8.398e-08 | -4.1664e-05 | -54.4905 | 0.10983 | 0 | 0.10983
-0.03 | 5.278e-10 |  | 39.2 | -0.03 | 9.831e-08 | -4.1566e-05 | -54.3619 | 0.12857 | 0 | 0.12857
-0.01999 | 6.16e-10 |  | 39.4 | -0.01999 | 1.1438e-07 | -4.1452e-05 | -54.2123 | 0.14959 | 0 | 0.14959
-0.009988 | 7.166e-10 |  | 39.6 | -0.009988 | 1.3326e-07 | -4.1318e-05 | -54.038 | 0.17428 | 0 | 0.17428
2.646e-05 | 8.223e-10 |  | 39.8 | 2.646e-05 | 1.5389e-07 | -4.1164e-05 | -53.8368 | 0.20126 | 0 | 0.20126
0.01001 | 9.48e-10 |  | 40 | 0.01001 | 1.7703e-07 | -4.0987e-05 | -53.6052 | 0.23153 | 0 | 0.23153
0.02003 | 1.082e-09 |  | 40.2 | 0.02003 | 2.03e-07 | -4.0784e-05 | -53.3398 | 0.26549 | 0 | 0.26549
0.03004 | 1.214e-09 |  | 40.4 | 0.03004 | 2.296e-07 | -4.0555e-05 | -53.0395 | 0.30028 | 0 | 0.30028
0.04004 | 1.355e-09 |  | 40.6 | 0.04004 | 2.569e-07 | -4.0298e-05 | -52.7035 | 0.33599 | 0 | 0.33599
0.05002 | 1.487e-09 |  | 40.8 | 0.05002 | 2.842e-07 | -4.0014e-05 | -52.3318 | 0.37169 | 0 | 0.37169
0.06001 | 1.595e-09 |  | 41 | 0.06001 | 3.082e-07 | -3.9705e-05 | -51.9287 | 0.40308 | 0 | 0.40308
0.07003 | 1.671e-09 |  | 41.2 | 0.07003 | 3.266e-07 | -3.9379e-05 | -51.5016 | 0.42714 | 0 | 0.42714
0.08001 | 1.743e-09 |  | 41.4 | 0.08001 | 3.414e-07 | -3.9037e-05 | -51.0551 | 0.4465 | 0 | 0.4465
0.09001 | 1.824e-09 |  | 41.6 | 0.09001 | 3.567e-07 | -3.8681e-05 | -50.5886 | 0.46651 | 0 | 0.46651
0.1001 | 1.919e-09 |  | 41.8 | 0.1001 | 3.743e-07 | -3.8306e-05 | -50.099 | 0.48953 | 0 | 0.48953
0.1101 | 2.004e-09 |  | 42 | 0.1101 | 3.923e-07 | -3.7914e-05 | -49.586 | 0.51307 | 0 | 0.51307
0.1201 | 2.097e-09 |  | 42.2 | 0.1201 | 4.101e-07 | -3.7504e-05 | -49.0496 | 0.53635 | 0 | 0.53635
0.1301 | 2.183e-09 |  | 42.4 | 0.1301 | 4.28e-07 | -3.7076e-05 | -48.4899 | 0.55976 | 0 | 0.55976
0.14 | 2.257e-09 |  | 42.6 | 0.14 | 4.44e-07 | -3.6632e-05 | -47.9092 | 0.58069 | 0 | 0.58069
0.1501 | 2.319e-09 |  | 42.8 | 0.1501 | 4.576e-07 | -3.6174e-05 | -47.3107 | 0.59847 | 0 | 0.59847
0.1601 | 2.374e-09 |  | 43 | 0.1601 | 4.693e-07 | -3.5705e-05 | -46.6969 | 0.61377 | 0 | 0.61377
0.17 | 2.422e-09 |  | 43.2 | 0.17 | 4.796e-07 | -3.5226e-05 | -46.0697 | 0.62724 | 0 | 0.62724
0.1801 | 2.453e-09 |  | 43.4 | 0.1801 | 4.875e-07 | -3.4738e-05 | -45.4321 | 0.63758 | 0 | 0.63758
0.19 | 2.469e-09 |  | 43.6 | 0.19 | 4.922e-07 | -3.4246e-05 | -44.7884 | 0.64372 | 0 | 0.64372
0.2 | 2.477e-09 |  | 43.8 | 0.2 | 4.946e-07 | -3.3751e-05 | -44.1415 | 0.64686 | 0 | 0.64686
0.21 | 2.468e-09 |  | 44 | 0.21 | 4.945e-07 | -3.3257e-05 | -43.4948 | 0.64673 | 0 | 0.64673
0.22 | 2.439e-09 |  | 44.2 | 0.22 | 4.907e-07 | -3.2766e-05 | -42.853 | 0.64176 | 0 | 0.64176
0.2301 | 2.394e-09 |  | 44.4 | 0.2301 | 4.833e-07 | -3.2283e-05 | -42.2209 | 0.63208 | 0 | 0.63208
0.24 | 2.337e-09 |  | 44.6 | 0.24 | 4.731e-07 | -3.181e-05 | -41.6022 | 0.61874 | 0 | 0.61874
0.25 | 2.266e-09 |  | 44.8 | 0.25 | 4.603e-07 | -3.1349e-05 | -41.0002 | 0.602 | 0 | 0.602
0.26 | 2.191e-09 |  | 45 | 0.26 | 4.457e-07 | -3.0904e-05 | -40.4173 | 0.58291 | 0 | 0.58291
0.27 | 2.097e-09 |  | 45.2 | 0.27 | 4.288e-07 | -3.0475e-05 | -39.8565 | 0.56081 | 0 | 0.56081
0.28 | 2.004e-09 |  | 45.4 | 0.28 | 4.101e-07 | -3.0065e-05 | -39.3201 | 0.53635 | 0 | 0.53635
0.2901 | 1.912e-09 |  | 45.6 | 0.2901 | 3.916e-07 | -2.9673e-05 | -38.808 | 0.51215 | 0 | 0.51215
0.3001 | 1.804e-09 |  | 45.8 | 0.3001 | 3.716e-07 | -2.9302e-05 | -38.322 | 0.486 | 0 | 0.486
0.3101 | 1.706e-09 |  | 46 | 0.3101 | 3.51e-07 | -2.8951e-05 | -37.8629 | 0.45906 | 0 | 0.45906
0.3201 | 1.615e-09 |  | 46.2 | 0.3201 | 3.321e-07 | -2.8618e-05 | -37.4286 | 0.43434 | 0 | 0.43434
0.3301 | 1.535e-09 |  | 46.4 | 0.3301 | 3.15e-07 | -2.8303e-05 | -37.0166 | 0.41197 | 0 | 0.41197
0.3401 | 1.46e-09 |  | 46.6 | 0.3401 | 2.995e-07 | -2.8004e-05 | -36.6249 | 0.3917 | 0 | 0.3917
0.3501 | 1.394e-09 |  | 46.8 | 0.3501 | 2.854e-07 | -2.7719e-05 | -36.2517 | 0.37326 | 0 | 0.37326
0.3601 | 1.347e-09 |  | 47 | 0.3601 | 2.741e-07 | -2.7444e-05 | -35.8932 | 0.35848 | 0 | 0.35848
0.3701 | 1.315e-09 |  | 47.2 | 0.3701 | 2.662e-07 | -2.7178e-05 | -35.545 | 0.34815 | 0 | 0.34815
0.3801 | 1.293e-09 |  | 47.4 | 0.3801 | 2.608e-07 | -2.6917e-05 | -35.2039 | 0.34109 | 0 | 0.34109
0.3901 | 1.285e-09 |  | 47.6 | 0.3901 | 2.578e-07 | -2.666e-05 | -34.8668 | 0.33716 | 0 | 0.33716
0.4001 | 1.294e-09 |  | 47.8 | 0.4001 | 2.579e-07 | -2.6402e-05 | -34.5295 | 0.33729 | 0 | 0.33729
0.4101 | 1.312e-09 |  | 48 | 0.4101 | 2.606e-07 | -2.6141e-05 | -34.1887 | 0.34083 | 0 | 0.34083
0.4201 | 1.335e-09 |  | 48.2 | 0.4201 | 2.647e-07 | -2.5876e-05 | -33.8425 | 0.34619 | 0 | 0.34619
0.4301 | 1.36e-09 |  | 48.4 | 0.4301 | 2.695e-07 | -2.5607e-05 | -33.49 | 0.35247 | 0 | 0.35247
0.4401 | 1.396e-09 |  | 48.6 | 0.4401 | 2.756e-07 | -2.5331e-05 | -33.1296 | 0.36044 | 0 | 0.36044
0.4501 | 1.436e-09 |  | 48.8 | 0.4501 | 2.832e-07 | -2.5048e-05 | -32.7592 | 0.37038 | 0 | 0.37038
0.4601 | 1.474e-09 |  | 49 | 0.4601 | 2.91e-07 | -2.4757e-05 | -32.3786 | 0.38058 | 0 | 0.38058
0.47 | 1.519e-09 |  | 49.2 | 0.47 | 2.993e-07 | -2.4458e-05 | -31.9872 | 0.39144 | 0 | 0.39144
0.4801 | 1.558e-09 |  | 49.4 | 0.4801 | 3.077e-07 | -2.415e-05 | -31.5847 | 0.40243 | 0 | 0.40243
0.4901 | 1.608e-09 |  | 49.6 | 0.4901 | 3.166e-07 | -2.3834e-05 | -31.1707 | 0.41407 | 0 | 0.41407
0.5001 | 1.639e-09 |  | 49.8 | 0.5001 | 3.247e-07 | -2.3509e-05 | -30.746 | 0.42466 | 0 | 0.42466
0.5101 | 1.673e-09 |  | 50 | 0.5101 | 3.312e-07 | -2.3178e-05 | -30.3128 | 0.43316 | 0 | 0.43316
0.5201 | 1.707e-09 |  | 50.2 | 0.5201 | 3.38e-07 | -2.284e-05 | -29.8708 | 0.44205 | 0 | 0.44205
0.5301 | 1.736e-09 |  | 50.4 | 0.5301 | 3.443e-07 | -2.2495e-05 | -29.4205 | 0.45029 | 0 | 0.45029
0.5401 | 1.756e-09 |  | 50.6 | 0.5401 | 3.492e-07 | -2.2146e-05 | -28.9638 | 0.4567 | 0 | 0.4567
0.5501 | 1.777e-09 |  | 50.8 | 0.5501 | 3.533e-07 | -2.1793e-05 | -28.5017 | 0.46206 | 0 | 0.46206
0.5601 | 1.792e-09 |  | 51 | 0.5601 | 3.569e-07 | -2.1436e-05 | -28.035 | 0.46677 | 0 | 0.46677
0.5701 | 1.796e-09 |  | 51.2 | 0.5701 | 3.588e-07 | -2.1077e-05 | -27.5657 | 0.46926 | 0 | 0.46926
0.5801 | 1.808e-09 |  | 51.4 | 0.5801 | 3.604e-07 | -2.0717e-05 | -27.0944 | 0.47135 | 0 | 0.47135
0.5901 | 1.813e-09 |  | 51.6 | 0.5901 | 3.621e-07 | -2.0355e-05 | -26.6208 | 0.47357 | 0 | 0.47357
0.6001 | 1.788e-09 |  | 51.8 | 0.6001 | 3.601e-07 | -1.9995e-05 | -26.1498 | 0.47096 | 0 | 0.47096
0.6101 | 1.777e-09 |  | 52 | 0.6101 | 3.565e-07 | -1.9638e-05 | -25.6836 | 0.46625 | 0 | 0.46625
0.6201 | 1.761e-09 |  | 52.2 | 0.6201 | 3.538e-07 | -1.9284e-05 | -25.2209 | 0.46272 | 0 | 0.46272
0.6301 | 1.743e-09 |  | 52.4 | 0.6301 | 3.504e-07 | -1.8934e-05 | -24.7626 | 0.45827 | 0 | 0.45827
0.6401 | 1.719e-09 |  | 52.6 | 0.6401 | 3.462e-07 | -1.8588e-05 | -24.3098 | 0.45278 | 0 | 0.45278
0.6501 | 1.696e-09 |  | 52.8 | 0.6501 | 3.415e-07 | -1.8246e-05 | -23.8632 | 0.44663 | 0 | 0.44663
0.6601 | 1.685e-09 |  | 53 | 0.6601 | 3.381e-07 | -1.7908e-05 | -23.421 | 0.44218 | 0 | 0.44218
0.6701 | 1.666e-09 |  | 53.2 | 0.6701 | 3.351e-07 | -1.7573e-05 | -22.9827 | 0.43826 | 0 | 0.43826
0.6801 | 1.66e-09 |  | 53.4 | 0.6801 | 3.326e-07 | -1.724e-05 | -22.5478 | 0.43499 | 0 | 0.43499
0.6901 | 1.646e-09 |  | 53.6 | 0.6901 | 3.306e-07 | -1.691e-05 | -22.1154 | 0.43237 | 0 | 0.43237
0.7001 | 1.641e-09 |  | 53.8 | 0.7001 | 3.287e-07 | -1.6581e-05 | -21.6855 | 0.42989 | 0 | 0.42989
0.7101 | 1.641e-09 |  | 54 | 0.7101 | 3.282e-07 | -1.6253e-05 | -21.2563 | 0.42924 | 0 | 0.42924
0.7201 | 1.646e-09 |  | 54.2 | 0.7201 | 3.287e-07 | -1.5924e-05 | -20.8264 | 0.42989 | 0 | 0.42989
0.7301 | 1.663e-09 |  | 54.4 | 0.7301 | 3.309e-07 | -1.5593e-05 | -20.3936 | 0.43277 | 0 | 0.43277
0.7401 | 1.688e-09 |  | 54.6 | 0.7401 | 3.351e-07 | -1.5258e-05 | -19.9553 | 0.43826 | 0 | 0.43826
0.7501 | 1.725e-09 |  | 54.8 | 0.7501 | 3.413e-07 | -1.4917e-05 | -19.509 | 0.44637 | 0 | 0.44637
0.7601 | 1.77e-09 |  | 55 | 0.7601 | 3.495e-07 | -1.4567e-05 | -19.0519 | 0.45709 | 0 | 0.45709
0.7701 | 1.831e-09 |  | 55.2 | 0.7701 | 3.601e-07 | -1.4207e-05 | -18.5809 | 0.47096 | 0 | 0.47096
0.7801 | 1.913e-09 |  | 55.4 | 0.7801 | 3.744e-07 | -1.3833e-05 | -18.0913 | 0.48966 | 0 | 0.48966
0.7901 | 2.014e-09 |  | 55.6 | 0.7901 | 3.927e-07 | -1.344e-05 | -17.5777 | 0.51359 | 0 | 0.51359
0.8001 | 2.177e-09 |  | 55.8 | 0.8001 | 4.191e-07 | -1.3021e-05 | -17.0295 | 0.54812 | 0 | 0.54812